## Method Comparison
Verification of NT-proBNP on Vitros 5600
Analyte: | NT-proBNP
Instrument: | Vitros 5600
Instrument S/N: | 1290353
Purpose: | This method serves to bringing in-house BNP testing for heart failure assessments
Unit of Measure: | ng/L
The complete evaluation consists of:
1) | Precision Studies over 20 days, Running Each Level Duplicates, Across 2 Runs; Total 80 points
2) | Accuracy by Method Comparison ( 20 - 40 samples minimum)
3) | Linearity (See Corresponding Tab)
4) | Reference Interval Verification (See Corresponding Tab) - only on recommendation from biochemist
Additional Studies Include:
Accuracy by Method Comparison
Test Method 1: | Vitros 5600
Principle
Reagent Lot #: | #032200
Calibrator Lot #: | #0132
Primary Reference Method: | Abbott Alinity i
Principle
Reagent Lot# (If available from Ref Lab)
Calibrator Lot# (If available from Ref Lab)
Instructions
Part 1
Please use 20 serum (SST) and 20 EDTA plasma samples
| a. | Please ensure the instrument is ready to report in SI units
| b. | Ensure the lower end of the AMR is set to 0
| b. | Ensure QC has passed for the day and instrument is in working order.
| c. | Run equal number of patient samples over the course of 5 days on the Test Method and Reference Method
| d. | Jot down the results table below.
| Note: | If calibrator or reagent lot number changes during course of study, please make note of this in the comment column
Matrix (U/S) | Sample No. | Sample ID | Primary Refence Method Result | Test Method Result | Tube Type (e.g. SST/EDTA) | Day | Comments
S | 1 | 13212 | 1485.1 | 1710 | SST | 1
S | 2 | 13213 | 855.2 | 881 | SST | 1
S | 3 | 13214 | 9964 | 10200 | SST | 1
S | 4 | 13215 | 13204.8 | 12700 | SST | 1
S | 5 | 13216 | 4948.7 | 4690 | SST | 1
S | 6 | 13217 | 11622.3 | 10900 | SST | 1
S | 7 | 13218 | 8807.8 | 7860 | SST | 1
S | 8 | 13219 | 2996.8 | 3150 | SST | 1
S | 9 | 13220 | 5692.5 | 5910 | SST | 1
S | 10 | 13221 | 2050.8 | 2310 | SST | 1
S | 11 | 13222 | 536.5 | 509 | SST | 1
S | 12 | 13223 | 5438.5 | 5890 | SST | 2
S | 13 | 13224 | 1658.9 | 1860 | SST | 2
S | 14 | 13225 | 5553.2 | 6310 | SST | 2
S | 15 | 13226 | 73.7 | 36.2 | SST | 2
S | 16 | 13227 | 3898.7 | 4250 | SST | 2
S | 17 | 13228 | 10846.7 | 10400 | SST | 2
S | 18 | 13229 | 8377.6 | 7380 | SST | 2
S | 19 | 13230 | 146.4 | 141 | SST | 2
S | 20 | 13231 | 2264.7 | 2580 | SST | 3
S | 21 | 13232 | 713.7 | 733 | SST | 3
S | 22 | 13233 | 1503.3 | 1640 | SST | 3
S | 23 | 13234 | 854.3 | 855 | SST | 3
S | 24 | 13235 | 1080.1 | 1240 | SST | 3
S | 25 | 13236 | 733.1 | 629 | SST | 3
S | 26 | 13237 | 5766.2 | 5390 | SST | 3
S | 27 | 13238 | 1530.4 | 1520 | SST | 4
S | 28 | 13239 | 636.4 | 674 | SST | 4
S | 29 | 13240 | 11958.1 | 13800 | SST | 4
S | 30 | 13241 | 1600 | 1780 | SST | 4
S | 31 | 13242 | 172.4 | 145 | SST | 4
S | 32 | 13243 | 5672.6 | 5710 | SST | 4
S | 33 | 13244 | 5825.9 | 5630 | SST | 4
S | 34 | 13245 | 344.1 | 358 | SST | 5
S | 35 | 13246 | 39.5 | 23.5 | SST | 5
S | 36 | 13247 | 250.4 | 256 | SST | 5
S | 37 | 13248 | 1002.4 | 1100 | SST | 5
S | 38 | 13249 | 3828.3 | 3470 | SST | 5
S | 39 | 13250 | 787.1 | 840 | SST | 5
S | 40 | 13251 | 2241.8 | 2310 | SST | 5

## Precision
Verification of NT-proBNP on Vitros 5600
Analytes: | NT-proBNP
Instrument: | Vitros 5600
Instrument S/N: | 1290353
Purpose: | This method serves to bringing in-house BNP testing for heart failure assessments
Matrix | Serum
Unit of Measure: | ng/L
Instructions for Precision Studies
Reagent Lot #: | 32200
Calibrator Lot #:
QC Lot Level 1
QC Lot Level 3
QC Lot Level 4
QC Lot Level 4 | N/A
a. | Perform assay calibration on the first day. For each day, run each level of QC across 2 runs, each run in duplicate.
b. | One run ideally in the morning (Run 1) with duplicate measurements & and another in PM (Run 2).
c. | Indicate any calibration. Ensure there is no change in QC Lot# throughout this study.
Note: | Please note calibrator, reagent, or QC lot number changes
| Level 1 | [5 empty] | Level 3 | [5 empty] | Level 4
| Day | Run | Replicate | Date | Result | Calibration(Y/N) | Day | Run | Replicate | Date | Result | Calibration(Y/N) | Day | Run | Replicate | Date | Result | Calibration(Y/N)
| 1 | 1 | 1 | 14/1/26 | 101 | N | 1 | 1 | 1 | 14/1/26 | 272 | N | 1 | 1 | 1 | 14/1/26 | 4640 | N
| 1 | 1 | 2 | 14/1/26 | 104 | N | 1 | 1 | 2 | 14/1/26 | 277 | N | 1 | 1 | 2 | 14/1/26 | 4690 | N
| 1 | 2 | 1 | 14/1/26 | 99.7 | N | 1 | 2 | 1 | 14/1/26 | 265 | N | 1 | 2 | 1 | 14/1/26 | 4580 | N
| 1 | 2 | 2 | 14/1/26 | 101 | N | 1 | 2 | 2 | 14/1/26 | 271 | N | 1 | 2 | 2 | 14/1/26 | 4680 | N
| 2 | 1 | 1 | 15/1/26 | 104 | N | 2 | 1 | 1 | 15/1/26 | 278 | N | 2 | 1 | 1 | 15/1/26 | 4610 | N
| 2 | 1 | 2 | 15/1/26 | 102 | N | 2 | 1 | 2 | 15/1/26 | 266 | N | 2 | 1 | 2 | 15/1/26 | 4700 | N
| 2 | 2 | 1 | 15/1/26 | 104 | N | 2 | 2 | 1 | 15/1/26 | 273 | N | 2 | 2 | 1 | 15/1/26 | 4690 | N
| 2 | 2 | 2 | 15/1/26 | 102 | N | 2 | 2 | 2 | 15/1/26 | 278 | N | 2 | 2 | 2 | 15/1/26 | 4710 | N
| 3 | 1 | 1 | 16/1/26 | 102 | N | 3 | 1 | 1 | 16/1/26 | 272 | N | 3 | 1 | 1 | 16/1/26 | 4710 | N
| 3 | 1 | 2 | 16/1/26 | 100 | N | 3 | 1 | 2 | 16/1/26 | 276 | N | 3 | 1 | 2 | 16/1/26 | 4710 | N
| 3 | 2 | 1 | 16/1/26 | 106 | N | 3 | 2 | 1 | 16/1/26 | 279 | N | 3 | 2 | 1 | 16/1/26 | 4730 | N
| 3 | 2 | 2 | 16/1/26 | 103 | N | 3 | 2 | 2 | 16/1/26 | 277 | N | 3 | 2 | 2 | 16/1/26 | 4630 | N
| 4 | 1 | 1 | 17/1/26 | 101 | N | 4 | 1 | 1 | 17/1/26 | 268 | N | 4 | 1 | 1 | 17/1/26 | 4610 | N
| 4 | 1 | 2 | 17/1/26 | 104 | N | 4 | 1 | 2 | 17/1/26 | 271 | N | 4 | 1 | 2 | 17/1/26 | 4720 | N
| 4 | 2 | 1 | 17/1/26 | 107 | N | 4 | 2 | 1 | 17/1/26 | 272 | N | 4 | 2 | 1 | 17/1/26 | 4750 | N
| 4 | 2 | 2 | 17/1/26 | 103 | N | 4 | 2 | 2 | 17/1/26 | 269 | N | 4 | 2 | 2 | 17/1/26 | 4530 | N
| 5 | 1 | 1 | 18/1/26 | 101 | N | 5 | 1 | 1 | 18/1/26 | 266 | N | 5 | 1 | 1 | 18/1/26 | 4680 | N
| 5 | 1 | 2 | 18/1/26 | 103 | N | 5 | 1 | 2 | 18/1/26 | 274 | N | 5 | 1 | 2 | 18/1/26 | 4790 | N
| 5 | 2 | 1 | 18/1/26 | 105 | N | 5 | 2 | 1 | 18/1/26 | 272 | N | 5 | 2 | 1 | 18/1/26 | 4740 | N
| 5 | 2 | 2 | 18/1/26 | 106 | N | 5 | 2 | 2 | 18/1/26 | 282 | N | 5 | 2 | 2 | 18/1/26 | 4800 | N
| 6 | 1 | 1 | 19/1/26 | 102 | N | 6 | 1 | 1 | 19/1/26 | 276 | N | 6 | 1 | 1 | 19/1/26 | 4520 | N
| 6 | 1 | 2 | 19/1/26 | 103 | N | 6 | 1 | 2 | 19/1/26 | 275 | N | 6 | 1 | 2 | 19/1/26 | 4720 | N
| 6 | 2 | 1 | 19/1/26 | 111 | N | 6 | 2 | 1 | 19/1/26 | 278 | N | 6 | 2 | 1 | 19/1/26 | 4550 | N
| 6 | 2 | 2 | 19/1/26 | 107 | N | 6 | 2 | 2 | 19/1/26 | 280 | N | 6 | 2 | 2 | 19/1/26 | 4750 | N
| 7 | 1 | 1 | 20/1/26 | 110 | N | 7 | 1 | 1 | 20/1/26 | 268 | N | 7 | 1 | 1 | 20/1/26 | 4700 | N
| 7 | 1 | 2 | 20/1/26 | 107 | N | 7 | 1 | 2 | 20/1/26 | 268 | N | 7 | 1 | 2 | 20/1/26 | 4590 | N
| 7 | 2 | 1 | 20/1/26 | 105 | N | 7 | 2 | 1 | 20/1/26 | 272 | N | 7 | 2 | 1 | 20/1/26 | 4720 | N
| 7 | 2 | 2 | 20/1/26 | 104 | N | 7 | 2 | 2 | 20/1/26 | 270 | N | 7 | 2 | 2 | 20/1/26 | 4730 | N
| 8 | 1 | 1 | 21/1/26 | 107 | N | 8 | 1 | 1 | 21/1/26 | 278 | N | 8 | 1 | 1 | 21/1/26 | 4850 | N
| 8 | 1 | 2 | 21/1/26 | 106 | N | 8 | 1 | 2 | 21/1/26 | 280 | N | 8 | 1 | 2 | 21/1/26 | 4760 | N
| 8 | 2 | 1 | 21/1/26 | 104 | N | 8 | 2 | 1 | 21/1/26 | 269 | N | 8 | 2 | 1 | 21/1/26 | 4830 | N
| 8 | 2 | 2 | 21/1/26 | 105 | N | 8 | 2 | 2 | 21/1/26 | 268 | N | 8 | 2 | 2 | 21/1/26 | 4720 | N
| 9 | 1 | 1 | 22/1/26 | 107 | N | 9 | 1 | 1 | 22/1/26 | 278 | N | 9 | 1 | 1 | 22/1/26 | 4700 | N
| 9 | 1 | 2 | 22/1/26 | 107 | N | 9 | 1 | 2 | 22/1/26 | 277 | N | 9 | 1 | 2 | 22/1/26 | 4840 | N
| 9 | 2 | 1 | 22/1/26 | 104 | N | 9 | 2 | 1 | 22/1/26 | 272 | N | 9 | 2 | 1 | 22/1/26 | 4720 | N
| 9 | 2 | 2 | 22/1/26 | 102 | N | 9 | 2 | 2 | 22/1/26 | 273 | N | 9 | 2 | 2 | 22/1/26 | 4830 | N
| 10 | 1 | 1 | 23/1/26 | 107 | N | 10 | 1 | 1 | 23/1/26 | 278 | N | 10 | 1 | 1 | 23/1/26 | 4740 | N
| 10 | 1 | 2 | 23/1/26 | 106 | N | 10 | 1 | 2 | 23/1/26 | 278 | N | 10 | 1 | 2 | 23/1/26 | 4710 | N
| 10 | 2 | 1 | 23/1/26 | 104 | N | 10 | 2 | 1 | 23/1/26 | 281 | N | 10 | 2 | 1 | 23/1/26 | 4750 | N
| 10 | 2 | 2 | 23/1/26 | 105 | N | 10 | 2 | 2 | 23/1/26 | 274 | N | 10 | 2 | 2 | 23/1/26 | 4710 | N
| 11 | 1 | 1 | 2026-11-02 | 102 | Y new lot# 032250 | 11 | 1 | 1 | 2026-11-02 | 269 | Y new lot# 032250 | 11 | 1 | 1 | 2026-11-02 | 4580 | Y new lot# 032250
| 11 | 1 | 2 | 2026-11-02 | 100 | Y new lot# 032250 | 11 | 1 | 2 | 2026-11-02 | 273 | Y new lot# 032250 | 11 | 1 | 2 | 2026-11-02 | 4710 | Y new lot# 032250
| 11 | 2 | 1 | 2026-12-02 | 99 | N | 11 | 2 | 1 | 2026-12-02 | 269 | N | 11 | 2 | 1 | 2026-12-02 | 4750 | N
| 11 | 2 | 2 | 2026-12-02 | 99 | N | 11 | 2 | 2 | 2026-12-02 | 270 | N | 11 | 2 | 2 | 2026-12-02 | 4690 | N
| 12 | 1 | 1 | 15/2/26 | 99 | N | 12 | 1 | 1 | 15/2/26 | 260 | N | 12 | 1 | 1 | 15/2/26 | 4640 | N
| 12 | 1 | 2 | 15/2/26 | 98.5 | N | 12 | 1 | 2 | 15/2/26 | 267 | N | 12 | 1 | 2 | 15/2/26 | 4630 | N
| 12 | 2 | 1 | 16/2/26 | 103 | N | 12 | 2 | 1 | 16/2/26 | 270 | N | 12 | 2 | 1 | 16/2/26 | 4650 | N
| 12 | 2 | 2 | 16/2/26 | 102 | N | 12 | 2 | 2 | 16/2/26 | 272 | N | 12 | 2 | 2 | 16/2/26 | 4740 | N
| 13 | 1 | 1 | 19/2/26 | 106 | N | 13 | 1 | 1 | 19/2/26 | 267 | N | 13 | 1 | 1 | 19/2/26 | 4630 | N
| 13 | 1 | 2 | 19/2/26 | 103 | N | 13 | 1 | 2 | 19/2/26 | 266 | N | 13 | 1 | 2 | 19/2/26 | 4720 | N
| 13 | 2 | 1 | 19/2/26 | 105 | N | 13 | 2 | 1 | 19/2/26 | 281 | N | 13 | 2 | 1 | 19/2/26 | 4800 | N
| 13 | 2 | 2 | 19/2/26 | 109 | N | 13 | 2 | 2 | 19/2/26 | 278 | N | 13 | 2 | 2 | 19/2/26 | 4750 | N
| 14 | 1 | 1 | 20/2/26 | 106 | N | 14 | 1 | 1 | 20/2/26 | 266 | N | 14 | 1 | 1 | 20/2/26 | 4740 | N
| 14 | 1 | 2 | 20/2/26 | 104 | N | 14 | 1 | 2 | 20/2/26 | 274 | N | 14 | 1 | 2 | 20/2/26 | 4790 | N
| 14 | 2 | 1 | 20/2/26 | 106 | N | 14 | 2 | 1 | 20/2/26 | 269 | N | 14 | 2 | 1 | 20/2/26 | 4810 | N
| 14 | 2 | 2 | 20/2/26 | 104 | N | 14 | 2 | 2 | 20/2/26 | 276 | N | 14 | 2 | 2 | 20/2/26 | 4710 | N
| 15 | 1 | 1 | 21/2/26 | 103 | N | 15 | 1 | 1 | 21/2/26 | 270 | N | 15 | 1 | 1 | 21/2/26 | 4640 | N
| 15 | 1 | 2 | 21/2/26 | 103 | N | 15 | 1 | 2 | 21/2/26 | 272 | N | 15 | 1 | 2 | 21/2/26 | 4750 | N
| 15 | 2 | 1 | 21/2/26 | 97.1 | N | 15 | 2 | 1 | 21/2/26 | 267 | N | 15 | 2 | 1 | 21/2/26 | 4730 | N
| 15 | 2 | 2 | 21/2/26 | 97.5 | N | 15 | 2 | 2 | 21/2/26 | 266 | N | 15 | 2 | 2 | 21/2/26 | 4680 | N
| 16 | 1 | 1 | 22/2/26 | 104 | N | 16 | 1 | 1 | 22/2/26 | 278 | N | 16 | 1 | 1 | 22/2/26 | 4860 | N
| 16 | 1 | 2 | 22/2/26 | 104 | N | 16 | 1 | 2 | 22/2/26 | 269 | N | 16 | 1 | 2 | 22/2/26 | 4810 | N
| 16 | 2 | 1 | 22/2/26 | 102 | N | 16 | 2 | 1 | 22/2/26 | 270 | N | 16 | 2 | 1 | 22/2/26 | 4750 | N
| 16 | 2 | 2 | 22/2/26 | 102 | N | 16 | 2 | 2 | 22/2/26 | 269 | N | 16 | 2 | 2 | 22/2/26 | 4730 | N
| 17 | 1 | 1 | 23/2/26 | 109 | N | 17 | 1 | 1 | 23/2/26 | 281 | N | 17 | 1 | 1 | 23/2/26 | 4820 | N
| 17 | 1 | 2 | 23/2/26 | 111 | N | 17 | 1 | 2 | 23/2/26 | 282 | N | 17 | 1 | 2 | 23/2/26 | 4810 | N
| 17 | 2 | 1 | 23/2/26 | 106 | N | 17 | 2 | 1 | 23/2/26 | 281 | N | 17 | 2 | 1 | 23/2/26 | 4860 | N
| 17 | 2 | 2 | 23/2/23 | 105 | N | 17 | 2 | 2 | 23/2/26 | 282 | N | 17 | 2 | 2 | 23/2/26 | 4860 | N
| 18 | 1 | 1 | 24/2/26 | 106 | N | 18 | 1 | 1 | 24/2/26 | 281 | N | 18 | 1 | 1 | 24/2/26 | 4860 | N
| 18 | 1 | 2 | 24/2/26 | 105 | N | 18 | 1 | 2 | 24/2/26 | 282 | N | 18 | 1 | 2 | 24/2/26 | 4860
| 18 | 2 | 1 | 24/2/26 | 105 | N | 18 | 2 | 1 | 24/2/26 | 272 | N | 18 | 2 | 1 | 24/2/26 | 4790 | N
| 18 | 2 | 2 | 24/2/26 | 102 | N | 18 | 2 | 2 | 24/2/26 | 278 | N | 18 | 2 | 2 | 24/2/26 | 4730 | N
| 19 | 1 | 1 | 25/2/26 | 103 | N | 19 | 1 | 1 | 25/2/26 | 274 | N | 19 | 1 | 1 | 25/2/26 | 4780 | N
| 19 | 1 | 2 | 25/2/26 | 107 | N | 19 | 1 | 2 | 25/2/26 | 274 | N | 19 | 1 | 2 | 25/2/26 | 4770 | N
| 19 | 2 | 1 | 25/2/26 | 102 | N | 19 | 2 | 1 | 25/2/26 | 280 | N | 19 | 2 | 1 | 25/2/26 | 4890 | N
| 19 | 2 | 2 | 25/2/26 | 106 | N | 19 | 2 | 2 | 25/2/26 | 274 | N | 19 | 2 | 2 | 25/2/26 | 4800 | N
| 20 | 1 | 1 | 26/2/26 | 101 | N | 20 | 1 | 1 | 26/2/26 | 270 | N | 20 | 1 | 1 | 26/2/26 | 4940 | N
| 20 | 1 | 2 | 26/2/26 | 102 | N | 20 | 1 | 2 | 26/2/26 | 273 | N | 20 | 1 | 2 | 26/2/26 | 4810 | N
| 20 | 2 | 1 | 26/2/26 | 106 | N | 20 | 2 | 1 | 26/2/26 | 283 | N | 20 | 2 | 1 | 26/2/26 | 4870 | N
| 20 | 2 | 2 | 26/2/26 | 108 | N | 20 | 2 | 2 | 26/2/26 | 281 | N | 20 | 2 | 2 | 26/2/26 | 4830 | N
| 21 | 1 | 1 |  |  |  | 21 | 1 | 1 |  |  |  | 21 | 1 | 1
| 21 | 1 | 2 |  |  |  | 21 | 1 | 2 |  |  |  | 21 | 1 | 2
| 21 | 2 | 1 |  |  |  | 21 | 2 | 1 |  |  |  | 21 | 2 | 1
| 21 | 2 | 2 |  |  |  | 21 | 2 | 2 |  |  |  | 21 | 2 | 2
| 22 | 1 | 1 |  |  |  | 22 | 1 | 1 |  |  |  | 22 | 1 | 1
| 22 | 1 | 2 |  |  |  | 22 | 1 | 2 |  |  |  | 22 | 1 | 2
| 22 | 2 | 1 |  |  |  | 22 | 2 | 1 |  |  |  | 22 | 2 | 1
| 22 | 2 | 2 |  |  |  | 22 | 2 | 2 |  |  |  | 22 | 2 | 2
| 23 | 1 | 1 |  |  |  | 23 | 1 | 1 |  |  |  | 23 | 1 | 1
| 23 | 1 | 2 |  |  |  | 23 | 1 | 2 |  |  |  | 23 | 1 | 2
| 23 | 2 | 1 |  |  |  | 23 | 2 | 1 |  |  |  | 23 | 2 | 1
| 23 | 2 | 2 |  |  |  | 23 | 2 | 2 |  |  |  | 23 | 2 | 2
| 24 | 1 | 1 |  |  |  | 24 | 1 | 1 |  |  |  | 24 | 1 | 1
| 24 | 1 | 2 |  |  |  | 24 | 1 | 2 |  |  |  | 24 | 1 | 2
| 24 | 2 | 1 |  |  |  | 24 | 2 | 1 |  |  |  | 24 | 2 | 1
| 24 | 2 | 2 |  |  |  | 24 | 2 | 2 |  |  |  | 24 | 2 | 2
| 25 | 1 | 1 |  |  |  | 25 | 1 | 1 |  |  |  | 25 | 1 | 1
| 25 | 1 | 2 |  |  |  | 25 | 1 | 2 |  |  |  | 25 | 1 | 2
| 25 | 2 | 1 |  |  |  | 25 | 2 | 1 |  |  |  | 25 | 2 | 1
| 25 | 2 | 2 |  |  |  | 25 | 2 | 2 |  |  |  | 25 | 2 | 2

## Linearity, Sr
Linearity Study
Unit of Measure: | nmol/L |  |  |  |  | Pass Criteria: Use ordinary least square regression, orthoglonal or non-parametric regression model to assess the correlation. Pearson R > 0.90 and expected differences should be within allowable performance limits (APL). Analytical Measurement Interval (AMI) verified is the minimum and maximum range of levels examined + error associated at the lower and upper limits. Allowable performance limits from: Review - Proficiency Testing Precision Goals and Allowable Performance Limits for Chemistry and Hematology. Revision date: 2022-05-10.
Matrix: | Serum
Package Insert AMI:
Extended AMI (from insert):
Name of Linearity Material:
Lot# :
Company (e.g. CAP, Maine):
NOTE: | If no linearity material is commercially available, one can use a high patient pool with dilution studies.
Instructions: | Follow manufacturer instructions. Generally expect 3 replicates per level.
If available, provide the expected mean and/or range from the manufacturer
Levels | Replicate | Measured Result | Average | Expected Result | Expected Range
Level 1 | 1 | 325.7 | 322.35 | 322.35
Level 1 | 2 | 319
Level 1 | 3
Level 2 | 1 | 273.5 | 272.8 | 290.115
Level 2 | 2 | 272.1
Level 2 | 3
Level 3 | 1 | 196.8 | 196.6 | 257.88
Level 3 | 2 | 196.4
Level 3 | 3
Level 4 | 1 | 169.9 | 169.5 | 225.65
Level 4 | 2 | 169.1
Level 4 | 3
Level 5 | 1 | 136.2 | 136.2 | 193.4
Level 5 | 2 | 135.8
Level 5 | 3
Level 6 | 1 | 116.1 | 115.6 | 161.18
Level 6 | 2 | 115.1
Level 6 | 3
Level 7 | 1 | 93.6 | 94.05 | 128.94
Level 7 | 2 | 94.5
Level 7 | 3
Level 8 | 1 | 71.6 | 71.3 | 96.71
Level 8 | 2 | 71
Level 8 | 3
Level 9 | 1 | 51.8 | 50.55 | 64.47
Level 9 | 2 | 49.3
Level 9 | 3
Level 10 | 1 | 27.8 | 27.7 | 32.235
Level 10 | 2 | 27.6
Level 10 | 3
Level 10 | 1 | 0 | 0 | 0
Level 10 | 2 | 0
Level 10 | 3

## Reference Interval
Verification of NT-proBNP on Vitros 5600
Analytes: | NT-proBNP
Instrument: | Vitros 5600
Instrument S/N: | 1290353
Purpose: | This method serves to bringing in-house BNP testing for heart failure assessments
Matrix:
Unit of Measure: | ng/L
Refernece Interval (RI) Adoption Method: |  |  | Selected Options (1-5):
1. Transference via Method Equivalence
2. Verification by 20 - 40 Ostensibily Health Individuals
3. Therapeutic interval from clinical input
4. Therapeutic interval from manufacturer
5. Other (Explain) |  |  | Adoption of Medical Decision Points from Clinical Cut-offs
Current RI:
Current Source of RI:
Proposed RI
Proposed Source of RI:
Healthy Individual # | Gender | Age | Result | Bacrode | Specimen#
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68
69
70
71
72
73
74
75
76
77
78