## Method Summary
| Verification Protocol and Data Collection Template for: |  |  |  |  | 2026-07-07
Creatinine on Atellica CI
Measurand: | Analyte: | Creatinine | Matrix: | Urine and Serum | Collection Container(s): | Serum, lithium heparin plasma, dipotassium EDTA plasma, sodium heparin plasma, urine
New Method Information: | Instrument: | Atellica CI | Manufacturer: | Siemens
| Instrument S/N: |  | Reagent Lot #: |  | Calibrator Lot #:
Purpose: | This method serves to verify the performance of Creatinine (Urine and Serum) on Atellica CI.
Reporting: | Patient Results: | Whole numbers | QC: | 1 decimal (x.x)
| Unit of Measure: | umol/L | Note: Urine creatinine maybe converted to mmol/L.
The complete evaluation consists of: | 1) Precision Studies over 20 days, Running Each Level Duplicates, Across 2 Runs; Total 80 points
2) Accuracy by Method Comparison ( 20 - 40 samples minimum)
3) Linearity (Verification of Analytical Measurement Interval, See Linearity Tab)
4) On recommendation from biochemist - Reference Interval Verification (See Corresponding Tab)
Additional studies (if any) include the following:
5)
6)
7)
8)
Hospital Site:
Performed by (MLT):
Date Submitted:

## Accuracy
Accuracy by Method Comparison Study
New Test Method: (Analyte + Instrument) | New Method: | Creatinine Atellica CI | Matrix: | Urine and Serum | Collection Container(s): | Serum, lithium heparin plasma, dipotassium EDTA plasma, sodium heparin plasma, urine
Primary Reference Method: (This is the current method that you are verifying/validating the new method against) | Current Method: | Creatinine2 Atellica CI | Matrix: | Urine | Collection Container(s):
Instructions
|  | Please use 40 (20 serum + 20 urine) samples that span the measurement interval (AMI) | [5 empty] | Pass Criteria: Use orthoglonal or non-parametric regression model to assess the correlation. Pearson R > 0.90 and median bias should be within allowable performance limits (APL). Allowable performance limits from: Review - Proficiency Testing Precision Goals and Allowable Performance Limits for Chemistry and Hematology. Revision date: 2022-05-10.
| a. | Please ensure the instrument is ready to report in SI units, regardless of how final reporting will be done.
| b. | Ensure the lower end of the AMI is set to 0
| b. | Ensure QC has passed for the day and instrument is in working order.
| c. | Run equal number of patient samples over the course of 5 days on the Test Method and Reference Method
| d. | Jot down the results table below. Note if matrix is serum (S) or urine (U)
| Note: | If calibrator or reagent lot number changes during course of study, please make note of this.
| Matrix (U/S/P/W) | Sample No. | Sample ID | Current Method | New Method | Tube Type (e.g. SST/PST) | Day
| S | 1 | 29503976 | 1.233 | 1.208 | PST | 1
| S | 2 | 29503975 | 1.089 | 1.043 | PST | 1
| S | 3 | 29503974 | 1.754 | 1.705 | PST | 1
| S | 4 | 29118636 | 2.315 | 2.265 | PST | 1
| S | 5 | 29118635 | 1.185 | 1.157 | PST | 1
| S | 6 | 29118634 | 1.087 | 1.035 | PST | 2
| S | 7 | 29118633 | 1.701 | 1.719 | PST | 2
| S | 8 | 29118632 | 0.74 | 0.729 | PST | 2
| S | 9 | 29118631 | 0.534 | 0.599 | PST | 2
| S | 10 | 29118630 | 2.197 | 2.145 | PST | 2
| S | 11 | 29028183 | 1.127 | 1.108 | PST | 3
| S | 12 | 29028182 | 0.994 | 1.004 | PST | 3
| S | 13 | 29028180 | 1.525 | 1.474 | PST | 3
| S | 14 | 29028179 | 2.06 | 1.991 | PST | 3
| S | 15 | 29028178 | 1.362 | 1.298 | PST | 3
| S | 16 | 29028177 | 1.218 | 1.26 | PST | 4
| S | 17 | 29028176 | 1.808 | 1.793 | PST | 4
| S | 18 | 29028175 | 1.278 | 1.238 | PST | 4
| S | 19 | 29028174 | 1.572 | 1.547 | PST | 4
| S | 20 | 29028173 | 1.037 | 1.041 | PST | 4
| S | 21 | 29028183 | 1.127 | 1.108 | PST | 5
| S | 22 | 29028182 | 0.994 | 1.004 | PST | 5
| S | 23 | 29028181 | 2.028 | 1.94 | PST | 5
| S | 24 | 29503977 | 1.181 | 1.083 | PST | 5
| S | 25 | 29503978 | 1.751 | 1.679 | PST | 5
| U | 26 |  |  |  |  | 5
| U | 27 |  |  |  |  | 5
| U | 28 |  |  |  |  | 5
| U | 29 |  |  |  |  | 5
| U | 30 |  |  |  |  | 5
| U | 31 |  |  |  |  | 5
| U | 32 |  |  |  |  | 5
| U | 33 |  |  |  |  | 5
| U | 34 |  |  |  |  | 5
| U | 35 |  |  |  |  | 5
| U | 36 |  |  |  |  | 5
| U | 37 |  |  |  |  | 5
| U | 38 |  |  |  |  | 5
| U | 39 |  |  |  |  | 5
| U | 40 |  |  |  |  | 5

## Precision
Precision Studies with QC Materials
Matrix | Urine and Serum
Unit of Measure: | umol/L
Instructions for Precision Studies
a. | For each level of QC please input the data in the tables below:
b. | For each day, run each level of QC across 2 runs, each run in duplicate.
One in the morning (Run 1) with duplicate measures & and another in PM (Run 2).
Note: | Please note calibrator or QC lot number changes or other concerns in the comment column
Reagent Lot #: | [5 empty] | Pass Criteria: Total CV or SD less than, or equal to precision goals from Review - Proficiency Testing Precision Goals and Allowable Performance Limits for Chemistry and Hematology. Revision date: 2022-05-10.
Calibrator Lot #:
QC Lot Level 1
QC Lot Level 2
QC Lot Level 3
QC Lot Level 4
| Level 1 - Serum | [6 empty] | Level 2 - Serum | [7 empty] | Level 1 - Urine | [6 empty] | Level 2 - Urine
Date | Day | Run | Replicate | Date | Result | QC Lot# | Comments | Day | Run | Replicate | Date | Result | QC Lot# | Comments | Date | Day | Run | Replicate | Date | Result | QC Lot# | Comments | Day | Run | Replicate | Date | Result | QC Lot# | Comments
| 0 | 1 | 1 |  |  |  |  | 0 | 1 | 1 | [5 empty] | 0 | 1 | 1 |  |  |  |  | 0 | 1 | 1
| 0 | 1 | 2 |  |  |  |  | 0 | 1 | 2 | [5 empty] | 0 | 1 | 2 |  |  |  |  | 0 | 1 | 2
| 0 | 2 | 1 |  |  |  |  | 0 | 2 | 1 | [5 empty] | 0 | 2 | 1 |  |  |  |  | 0 | 2 | 1
| 0 | 2 | 2 |  |  |  |  | 0 | 2 | 2 | [5 empty] | 0 | 2 | 2 |  |  |  |  | 0 | 2 | 2
| 1 | 1 | 1 |  |  |  |  | 1 | 1 | 1 | [5 empty] | 1 | 1 | 1 |  |  |  |  | 1 | 1 | 1
| 1 | 1 | 2 |  |  |  |  | 1 | 1 | 2 | [5 empty] | 1 | 1 | 2 |  |  |  |  | 1 | 1 | 2
| 1 | 2 | 1 |  |  |  |  | 1 | 2 | 1 | [5 empty] | 1 | 2 | 1 |  |  |  |  | 1 | 2 | 1
| 1 | 2 | 2 |  |  |  |  | 1 | 2 | 2 | [5 empty] | 1 | 2 | 2 |  |  |  |  | 1 | 2 | 2
| 2 | 1 | 1 |  |  |  |  | 2 | 1 | 1 | [5 empty] | 2 | 1 | 1 |  |  |  |  | 2 | 1 | 1
| 2 | 1 | 2 |  |  |  |  | 2 | 1 | 2 | [5 empty] | 2 | 1 | 2 |  |  |  |  | 2 | 1 | 2
| 2 | 2 | 1 |  |  |  |  | 2 | 2 | 1 | [5 empty] | 2 | 2 | 1 |  |  |  |  | 2 | 2 | 1
| 2 | 2 | 2 |  |  |  |  | 2 | 2 | 2 | [5 empty] | 2 | 2 | 2 |  |  |  |  | 2 | 2 | 2
| 3 | 1 | 1 |  |  |  |  | 3 | 1 | 1 | [5 empty] | 3 | 1 | 1 |  |  |  |  | 3 | 1 | 1
| 3 | 1 | 2 |  |  |  |  | 3 | 1 | 2 | [5 empty] | 3 | 1 | 2 |  |  |  |  | 3 | 1 | 2
| 3 | 2 | 1 |  |  |  |  | 3 | 2 | 1 | [5 empty] | 3 | 2 | 1 |  |  |  |  | 3 | 2 | 1
| 3 | 2 | 2 |  |  |  |  | 3 | 2 | 2 | [5 empty] | 3 | 2 | 2 |  |  |  |  | 3 | 2 | 2
| 4 | 1 | 1 |  |  |  |  | 4 | 1 | 1 | [5 empty] | 4 | 1 | 1 |  |  |  |  | 4 | 1 | 1
| 4 | 1 | 2 |  |  |  |  | 4 | 1 | 2 | [5 empty] | 4 | 1 | 2 |  |  |  |  | 4 | 1 | 2
| 4 | 2 | 1 |  |  |  |  | 4 | 2 | 1 | [5 empty] | 4 | 2 | 1 |  |  |  |  | 4 | 2 | 1
| 4 | 2 | 2 |  |  |  |  | 4 | 2 | 2 | [5 empty] | 4 | 2 | 2 |  |  |  |  | 4 | 2 | 2
| 5 | 1 | 1 |  |  |  |  | 5 | 1 | 1 | [5 empty] | 5 | 1 | 1 |  |  |  |  | 5 | 1 | 1
| 5 | 1 | 2 |  |  |  |  | 5 | 1 | 2 | [5 empty] | 5 | 1 | 2 |  |  |  |  | 5 | 1 | 2
| 5 | 2 | 1 |  |  |  |  | 5 | 2 | 1 | [5 empty] | 5 | 2 | 1 |  |  |  |  | 5 | 2 | 1
| 5 | 2 | 2 |  |  |  |  | 5 | 2 | 2 | [5 empty] | 5 | 2 | 2 |  |  |  |  | 5 | 2 | 2
| 6 | 1 | 1 |  |  |  |  | 6 | 1 | 1 | [5 empty] | 6 | 1 | 1 |  |  |  |  | 6 | 1 | 1
| 6 | 1 | 2 |  |  |  |  | 6 | 1 | 2 | [5 empty] | 6 | 1 | 2 |  |  |  |  | 6 | 1 | 2
| 6 | 2 | 1 |  |  |  |  | 6 | 2 | 1 | [5 empty] | 6 | 2 | 1 |  |  |  |  | 6 | 2 | 1
| 6 | 2 | 2 |  |  |  |  | 6 | 2 | 2 | [5 empty] | 6 | 2 | 2 |  |  |  |  | 6 | 2 | 2
| 7 | 1 | 1 |  |  |  |  | 7 | 1 | 1 | [5 empty] | 7 | 1 | 1 |  |  |  |  | 7 | 1 | 1
| 7 | 1 | 2 |  |  |  |  | 7 | 1 | 2 | [5 empty] | 7 | 1 | 2 |  |  |  |  | 7 | 1 | 2
| 7 | 2 | 1 |  |  |  |  | 7 | 2 | 1 | [5 empty] | 7 | 2 | 1 |  |  |  |  | 7 | 2 | 1
| 7 | 2 | 2 |  |  |  |  | 7 | 2 | 2 | [5 empty] | 7 | 2 | 2 |  |  |  |  | 7 | 2 | 2
| 8 | 1 | 1 |  |  |  |  | 8 | 1 | 1 | [5 empty] | 8 | 1 | 1 |  |  |  |  | 8 | 1 | 1
| 8 | 1 | 2 |  |  |  |  | 8 | 1 | 2 | [5 empty] | 8 | 1 | 2 |  |  |  |  | 8 | 1 | 2
| 8 | 2 | 1 |  |  |  |  | 8 | 2 | 1 | [5 empty] | 8 | 2 | 1 |  |  |  |  | 8 | 2 | 1
| 8 | 2 | 2 |  |  |  |  | 8 | 2 | 2 | [5 empty] | 8 | 2 | 2 |  |  |  |  | 8 | 2 | 2
| 9 | 1 | 1 |  |  |  |  | 9 | 1 | 1 | [5 empty] | 9 | 1 | 1 |  |  |  |  | 9 | 1 | 1
| 9 | 1 | 2 |  |  |  |  | 9 | 1 | 2 | [5 empty] | 9 | 1 | 2 |  |  |  |  | 9 | 1 | 2
| 9 | 2 | 1 |  |  |  |  | 9 | 2 | 1 | [5 empty] | 9 | 2 | 1 |  |  |  |  | 9 | 2 | 1
| 9 | 2 | 2 |  |  |  |  | 9 | 2 | 2 | [5 empty] | 9 | 2 | 2 |  |  |  |  | 9 | 2 | 2
| 10 | 1 | 1 |  |  |  |  | 10 | 1 | 1 | [5 empty] | 10 | 1 | 1 |  |  |  |  | 10 | 1 | 1
| 10 | 1 | 2 |  |  |  |  | 10 | 1 | 2 | [5 empty] | 10 | 1 | 2 |  |  |  |  | 10 | 1 | 2
| 10 | 2 | 1 |  |  |  |  | 10 | 2 | 1 | [5 empty] | 10 | 2 | 1 |  |  |  |  | 10 | 2 | 1
| 10 | 2 | 2 |  |  |  |  | 10 | 2 | 2 | [5 empty] | 10 | 2 | 2 |  |  |  |  | 10 | 2 | 2
| 11 | 1 | 1 |  |  |  |  | 11 | 1 | 1 | [5 empty] | 11 | 1 | 1 |  |  |  |  | 11 | 1 | 1
| 11 | 1 | 2 |  |  |  |  | 11 | 1 | 2 | [5 empty] | 11 | 1 | 2 |  |  |  |  | 11 | 1 | 2
| 11 | 2 | 1 |  |  |  |  | 11 | 2 | 1 | [5 empty] | 11 | 2 | 1 |  |  |  |  | 11 | 2 | 1
| 11 | 2 | 2 |  |  |  |  | 11 | 2 | 2 | [5 empty] | 11 | 2 | 2 |  |  |  |  | 11 | 2 | 2
| 12 | 1 | 1 |  |  |  |  | 12 | 1 | 1 | [5 empty] | 12 | 1 | 1 |  |  |  |  | 12 | 1 | 1
| 12 | 1 | 2 |  |  |  |  | 12 | 1 | 2 | [5 empty] | 12 | 1 | 2 |  |  |  |  | 12 | 1 | 2
| 12 | 2 | 1 |  |  |  |  | 12 | 2 | 1 | [5 empty] | 12 | 2 | 1 |  |  |  |  | 12 | 2 | 1
| 12 | 2 | 2 |  |  |  |  | 12 | 2 | 2 | [5 empty] | 12 | 2 | 2 |  |  |  |  | 12 | 2 | 2
| 13 | 1 | 1 |  |  |  |  | 13 | 1 | 1 | [5 empty] | 13 | 1 | 1 |  |  |  |  | 13 | 1 | 1
| 13 | 1 | 2 |  |  |  |  | 13 | 1 | 2 | [5 empty] | 13 | 1 | 2 |  |  |  |  | 13 | 1 | 2
| 13 | 2 | 1 |  |  |  |  | 13 | 2 | 1 | [5 empty] | 13 | 2 | 1 |  |  |  |  | 13 | 2 | 1
| 13 | 2 | 2 |  |  |  |  | 13 | 2 | 2 | [5 empty] | 13 | 2 | 2 |  |  |  |  | 13 | 2 | 2
| 14 | 1 | 1 |  |  |  |  | 14 | 1 | 1 | [5 empty] | 14 | 1 | 1 |  |  |  |  | 14 | 1 | 1
| 14 | 1 | 2 |  |  |  |  | 14 | 1 | 2 | [5 empty] | 14 | 1 | 2 |  |  |  |  | 14 | 1 | 2
| 14 | 2 | 1 |  |  |  |  | 14 | 2 | 1 | [5 empty] | 14 | 2 | 1 |  |  |  |  | 14 | 2 | 1
| 14 | 2 | 2 |  |  |  |  | 14 | 2 | 2 | [5 empty] | 14 | 2 | 2 |  |  |  |  | 14 | 2 | 2
| 15 | 1 | 1 |  |  |  |  | 15 | 1 | 1 | [5 empty] | 15 | 1 | 1 |  |  |  |  | 15 | 1 | 1
| 15 | 1 | 2 |  |  |  |  | 15 | 1 | 2 | [5 empty] | 15 | 1 | 2 |  |  |  |  | 15 | 1 | 2
| 15 | 2 | 1 |  |  |  |  | 15 | 2 | 1 | [5 empty] | 15 | 2 | 1 |  |  |  |  | 15 | 2 | 1
| 15 | 2 | 2 |  |  |  |  | 15 | 2 | 2 | [5 empty] | 15 | 2 | 2 |  |  |  |  | 15 | 2 | 2
| 16 | 1 | 1 |  |  |  |  | 16 | 1 | 1 | [5 empty] | 16 | 1 | 1 |  |  |  |  | 16 | 1 | 1
| 16 | 1 | 2 |  |  |  |  | 16 | 1 | 2 | [5 empty] | 16 | 1 | 2 |  |  |  |  | 16 | 1 | 2
| 16 | 2 | 1 |  |  |  |  | 16 | 2 | 1 | [5 empty] | 16 | 2 | 1 |  |  |  |  | 16 | 2 | 1
| 16 | 2 | 2 |  |  |  |  | 16 | 2 | 2 | [5 empty] | 16 | 2 | 2 |  |  |  |  | 16 | 2 | 2
| 17 | 1 | 1 |  |  |  |  | 17 | 1 | 1 | [5 empty] | 17 | 1 | 1 |  |  |  |  | 17 | 1 | 1
| 17 | 1 | 2 |  |  |  |  | 17 | 1 | 2 | [5 empty] | 17 | 1 | 2 |  |  |  |  | 17 | 1 | 2
| 17 | 2 | 1 |  |  |  |  | 17 | 2 | 1 | [5 empty] | 17 | 2 | 1 |  |  |  |  | 17 | 2 | 1
| 17 | 2 | 2 |  |  |  |  | 17 | 2 | 2 | [5 empty] | 17 | 2 | 2 |  |  |  |  | 17 | 2 | 2
| 18 | 1 | 1 |  |  |  |  | 18 | 1 | 1 | [5 empty] | 18 | 1 | 1 |  |  |  |  | 18 | 1 | 1
| 18 | 1 | 2 |  |  |  |  | 18 | 1 | 2 | [5 empty] | 18 | 1 | 2 |  |  |  |  | 18 | 1 | 2
| 18 | 2 | 1 |  |  |  |  | 18 | 2 | 1 | [5 empty] | 18 | 2 | 1 |  |  |  |  | 18 | 2 | 1
| 18 | 2 | 2 |  |  |  |  | 18 | 2 | 2 | [5 empty] | 18 | 2 | 2 |  |  |  |  | 18 | 2 | 2
| 19 | 1 | 1 |  |  |  |  | 19 | 1 | 1 | [5 empty] | 19 | 1 | 1 |  |  |  |  | 19 | 1 | 1
| 19 | 1 | 2 |  |  |  |  | 19 | 1 | 2 | [5 empty] | 19 | 1 | 2 |  |  |  |  | 19 | 1 | 2
| 19 | 2 | 1 |  |  |  |  | 19 | 2 | 1 | [5 empty] | 19 | 2 | 1 |  |  |  |  | 19 | 2 | 1
| 19 | 2 | 2 |  |  |  |  | 19 | 2 | 2 | [5 empty] | 19 | 2 | 2 |  |  |  |  | 19 | 2 | 2
| 20 | 1 | 1 |  |  |  |  | 20 | 1 | 1 | [5 empty] | 20 | 1 | 1 |  |  |  |  | 20 | 1 | 1
| 20 | 1 | 2 |  |  |  |  | 20 | 1 | 2 | [5 empty] | 20 | 1 | 2 |  |  |  |  | 20 | 1 | 2
| 20 | 2 | 1 |  |  |  |  | 20 | 2 | 1 | [5 empty] | 20 | 2 | 1 |  |  |  |  | 20 | 2 | 1
| 20 | 2 | 2 |  |  |  |  | 20 | 2 | 2 | [5 empty] | 20 | 2 | 2 |  |  |  |  | 20 | 2 | 2
| 21 | 1 | 1 |  |  |  |  | 21 | 1 | 1 | [5 empty] | 21 | 1 | 1 |  |  |  |  | 21 | 1 | 1
| 21 | 1 | 2 |  |  |  |  | 21 | 1 | 2 | [5 empty] | 21 | 1 | 2 |  |  |  |  | 21 | 1 | 2
| 21 | 2 | 1 |  |  |  |  | 21 | 2 | 1 | [5 empty] | 21 | 2 | 1 |  |  |  |  | 21 | 2 | 1
| 21 | 2 | 2 |  |  |  |  | 21 | 2 | 2 | [5 empty] | 21 | 2 | 2 |  |  |  |  | 21 | 2 | 2
| 22 | 1 | 1 |  |  |  |  | 22 | 1 | 1 | [5 empty] | 22 | 1 | 1 |  |  |  |  | 22 | 1 | 1
| 22 | 1 | 2 |  |  |  |  | 22 | 1 | 2 | [5 empty] | 22 | 1 | 2 |  |  |  |  | 22 | 1 | 2
| 22 | 2 | 1 |  |  |  |  | 22 | 2 | 1 | [5 empty] | 22 | 2 | 1 |  |  |  |  | 22 | 2 | 1
| 22 | 2 | 2 |  |  |  |  | 22 | 2 | 2 | [5 empty] | 22 | 2 | 2 |  |  |  |  | 22 | 2 | 2
| 23 | 1 | 1 |  |  |  |  | 23 | 1 | 1 | [5 empty] | 23 | 1 | 1 |  |  |  |  | 23 | 1 | 1
| 23 | 1 | 2 |  |  |  |  | 23 | 1 | 2 | [5 empty] | 23 | 1 | 2 |  |  |  |  | 23 | 1 | 2
| 23 | 2 | 1 |  |  |  |  | 23 | 2 | 1 | [5 empty] | 23 | 2 | 1 |  |  |  |  | 23 | 2 | 1
| 23 | 2 | 2 |  |  |  |  | 23 | 2 | 2 | [5 empty] | 23 | 2 | 2 |  |  |  |  | 23 | 2 | 2
| 24 | 1 | 1 |  |  |  |  | 24 | 1 | 1 | [5 empty] | 24 | 1 | 1 |  |  |  |  | 24 | 1 | 1
| 24 | 1 | 2 |  |  |  |  | 24 | 1 | 2 | [5 empty] | 24 | 1 | 2 |  |  |  |  | 24 | 1 | 2
| 24 | 2 | 1 |  |  |  |  | 24 | 2 | 1 | [5 empty] | 24 | 2 | 1 |  |  |  |  | 24 | 2 | 1
| 24 | 2 | 2 |  |  |  |  | 24 | 2 | 2 | [5 empty] | 24 | 2 | 2 |  |  |  |  | 24 | 2 | 2
| 25 | 1 | 1 |  |  |  |  | 25 | 1 | 1 | [5 empty] | 25 | 1 | 1 |  |  |  |  | 25 | 1 | 1
| 25 | 1 | 2 |  |  |  |  | 25 | 1 | 2 | [5 empty] | 25 | 1 | 2 |  |  |  |  | 25 | 1 | 2
| 25 | 2 | 1 |  |  |  |  | 25 | 2 | 1 | [5 empty] | 25 | 2 | 1 |  |  |  |  | 25 | 2 | 1
| 25 | 2 | 2 |  |  |  |  | 25 | 2 | 2 | [5 empty] | 25 | 2 | 2 |  |  |  |  | 25 | 2 | 2
| 26 | 1 | 1 |  |  |  |  | 26 | 1 | 1 | [5 empty] | 26 | 1 | 1 |  |  |  |  | 26 | 1 | 1
| 26 | 1 | 2 |  |  |  |  | 26 | 1 | 2 | [5 empty] | 26 | 1 | 2 |  |  |  |  | 26 | 1 | 2
| 26 | 2 | 1 |  |  |  |  | 26 | 2 | 1 | [5 empty] | 26 | 2 | 1 |  |  |  |  | 26 | 2 | 1
| 26 | 2 | 2 |  |  |  |  | 26 | 2 | 2 | [5 empty] | 26 | 2 | 2 |  |  |  |  | 26 | 2 | 2
| 27 | 1 | 1 |  |  |  |  | 27 | 1 | 1 | [5 empty] | 27 | 1 | 1 |  |  |  |  | 27 | 1 | 1
| 27 | 1 | 2 |  |  |  |  | 27 | 1 | 2 | [5 empty] | 27 | 1 | 2 |  |  |  |  | 27 | 1 | 2
| 27 | 2 | 1 |  |  |  |  | 27 | 2 | 1 | [5 empty] | 27 | 2 | 1 |  |  |  |  | 27 | 2 | 1
| 27 | 2 | 2 |  |  |  |  | 27 | 2 | 2 | [5 empty] | 27 | 2 | 2 |  |  |  |  | 27 | 2 | 2
| 28 | 1 | 1 |  |  |  |  | 28 | 1 | 1 | [5 empty] | 28 | 1 | 1 |  |  |  |  | 28 | 1 | 1
| 28 | 1 | 2 |  |  |  |  | 28 | 1 | 2 | [5 empty] | 28 | 1 | 2 |  |  |  |  | 28 | 1 | 2
| 28 | 2 | 1 |  |  |  |  | 28 | 2 | 1 | [5 empty] | 28 | 2 | 1 |  |  |  |  | 28 | 2 | 1
| 28 | 2 | 2 |  |  |  |  | 28 | 2 | 2 | [5 empty] | 28 | 2 | 2 |  |  |  |  | 28 | 2 | 2
| 29 | 1 | 1 |  |  |  |  | 29 | 1 | 1 | [5 empty] | 29 | 1 | 1 |  |  |  |  | 29 | 1 | 1
| 29 | 1 | 2 |  |  |  |  | 29 | 1 | 2 | [5 empty] | 29 | 1 | 2 |  |  |  |  | 29 | 1 | 2
| 29 | 2 | 1 |  |  |  |  | 29 | 2 | 1 | [5 empty] | 29 | 2 | 1 |  |  |  |  | 29 | 2 | 1
| 29 | 2 | 2 |  |  |  |  | 29 | 2 | 2 | [5 empty] | 29 | 2 | 2 |  |  |  |  | 29 | 2 | 2

## Linearity, Sr
Linearity Study
Unit of Measure: | umol/L |  |  |  |  | Pass Criteria: Use ordinary least square regression, orthoglonal or non-parametric regression model to assess the correlation. Pearson R > 0.90 and expected differences should be within allowable performance limits (APL). Analytical Measurement Interval (AMI) verified is the minimum and maximum range of levels examined + error associated at the lower and upper limits. Allowable performance limits from: Review - Proficiency Testing Precision Goals and Allowable Performance Limits for Chemistry and Hematology. Revision date: 2022-05-10.
Matrix: | Serum
Package Insert AMI:
Extended AMI (from insert):
Name of Linearity Material:
Lot# :
Company (e.g. CAP, Maine):
NOTE: | If no linearity material is commercially available, one can use a high patient pool with dilution studies.
Instructions: | Follow manufacturer instructions. Generally expect 3 replicates per level.
If available, provide the expected mean and/or range from the manufacturer
Levels | Replicate | Measured Result | Average | Expected Result | Expected Range
Level 1 | 1 |  | #DIV/0!
Level 1 | 2
Level 1 | 3
Level 2 | 1 |  | #DIV/0!
Level 2 | 2
Level 2 | 3
Level 3 | 1 |  | #DIV/0!
Level 3 | 2
Level 3 | 3
Level 4 | 1 |  | #DIV/0!
Level 4 | 2
Level 4 | 3
Level 5 | 1 |  | #DIV/0!
Level 5 | 2
Level 5 | 3

## Linearity, Ur
Linearity Study
Unit of Measure: | umol/L |  |  |  |  | Pass Criteria: Use ordinary least square regression, orthoglonal or non-parametric regression model to assess the correlation. Pearson R > 0.90 and expected differences should be within allowable performance limits (APL). Analytical Measurement Interval (AMI) verified is the minimum and maximum range of levels examined + error associated at the lower and upper limits. Allowable performance limits from: Review - Proficiency Testing Precision Goals and Allowable Performance Limits for Chemistry and Hematology. Revision date: 2022-05-10.
Matrix: | Urine
Package Insert AMI:
Extended AMI (from insert):
Name of Linearity Material:
Lot# :
Company (e.g. CAP, Maine):
NOTE: | If no linearity material is commercially available, one can use a high patient pool with dilution studies.
Instructions: | Follow manufacturer instructions. Generally expect 3 replicates per level.
If available, provide the expected mean and/or range from the manufacturer
Levels | Replicate | Measured Result | Average | Expected Result | Expected Range
Level 1 | 1 |  | #DIV/0!
Level 1 | 2
Level 1 | 3
Level 2 | 1 |  | #DIV/0!
Level 2 | 2
Level 2 | 3
Level 3 | 1 |  | #DIV/0!
Level 3 | 2
Level 3 | 3
Level 4 | 1 |  | #DIV/0!
Level 4 | 2
Level 4 | 3
Level 5 | 1 |  | #DIV/0!
Level 5 | 2
Level 5 | 3

## Reference Interval
Reference Interval Verification
Matrix: | Urine and Serum
Unit of Measure: | umol/L
Current RI:
Current Source of RI:
Proposed RI:
Proposed Source of RI:
Refernece Interval (RI) Verification Method:
1. Transference via Method Equivalence
2. Verification by 20 - 40 Ostensibily Health Individuals
3. Therapeutic interval from clinical input
4. Therapeutic interval from manufacturer
5. Other (Explain)
Pass Criteria: At least 90% of the reference individuals are within the proposed reference limits in the absence of method equivalence.
| Healthy Individual # | Gender (M/F) | Age | Result | Bacrode | Other ID
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40

## Carryover
Verification of Creatinine on Atellica CI
Analytes: | Creatinine
Instrument: | Atellica CI
Instrument S/N: | 0
Unit of Measure: | umol/L
Matrix: | Urine and Serum
Package Insert AMI:
Extended AMI (from insert):
Name of Linearity Material:
Lot# :
Company (e.g. CAP, Maine):
If no linearity material is available commercially, one can use a high patient pool.
Sample (Urine/serum) Carryover Instructions
a. Prepare High Pool (ideally urine sample).
b. Prepare Low sample Pool (ideal serum matrix <110 umol/L), this will be called sampled L.
c. Load the samples in the order listed below and record results as indicated below.
| # | Sample | Result | Low-Low | High-Low
| 1 | L | 255
| 2 | L | 255
| 3 | L | 251
| 4 | H | 764
| 5 | H | 762
| 6 | L | 252
| 7 | L | 255
| 8 | H | 765
| 9 | H | 765
| 10 | L | 254
| 11 | L | 253
| 12 | L | 254
| 13 | L | 255
| 14 | H | 766
| 15 | H | 763
| 16 | H | 765
| 17 | H | 765
| 18 | L | 254
| 19 | L | 252
| 20 | L | 254
Mean
SD
Carry-Over